Health and Safety hazards and control measures

Hazardous chemical - Xylene

Xylene is a Flammable. Irritating to the skin, eyes, mucous membranes, and respiratory tract. A carcinogen that may damage the liver and kidneys through prolonged/repeated exposure. Acute toxicity.  Wear either nitrile gloves or xylene resistant gloves.  Always use in a chemical fume hood.  Do not wash into drains, dispose via chemical waste route.  Store in fume cabinet separate to flammable solvents.  Note: if nitrile gloves are used, use forceps to minimise exposure to xylene.

Hazardous chemical - Ethanol

Ethanol is a highly flammable liquid and vapour.   It causes serious eye irritation.  Wear nitrile gloves.  Low risk of eye splash as ethanol steps are performed in a fume hood.  Avoid use near flames, heat or sparks.

Hazardous chemical – Triton X-100

Triton X-100 is a corrosive liquid. It can case mild skin irritation, serious eye damage and may cause respiratory irritation. Wear either nitrile gloves or xylene resistant gloves.  Always use in a chemical fume hood.  Do not wash into drains, dispose via chemical waste route.

Steamer – Hot steam may cause burns.  Heat may cause glass containers to break.  Staff/students are trained in the safe use of the steamer. Only use plastic containers in the steamer.

PROTOCOL

Imaging Mass Cytometry

- All tissue stained for the IMC should be between 5-10µm thick to ensure optimal results.

- Fluidigm recommends conjugation of antibodies with their MaxPar X8 Kit

- Regions of Interest on the tissue are drawn and then ablated using a laser at a frequency of 200Hz at a 1μm resolution

- Ensure all channel ‘.tiffs’ and their associated masks per sample are stored in a distinct folder or the HistoCAT software won’t open the images as needed

Staining

- Dewax slides in a Xylene (Honeywell) bath for 5 minutes, then move the slides into a second Xylene bath for another 5 minutes – must be carried out in a fume hood

- Rehydrate the slides through an ethanol (VWR) series, 100%, 100%, 90%, 70% for five minutes each, before cleansing in dH2O – must be carried out in a fume hood

- Perform an antigen retrieval in a pH9 1xEDTA (30ml 10XEDTA/270ml dH2O) buffer, using a steamer (CookWorks) for 20 minutes

- Cool the slides in ice, then cleanse in dH20 followed by 1xPBS for 5 minutes each

- Draw a hydrophobic barrier around the tissue with a PAP pen (abcam) and treat with 10% normal horse serum (Vector) in primary diluent for 45 minutes

- Treat the slides with a pre-prepared antibody cocktail, made of your conjugates diluted in primary diluent, then incubate at 4°C overnight

- Treat the slides in 0.02% Triton X-100 (Sigma-Aldrich) in dH2O for 8 minutes, the slides are then cleansed in dH20 twice for 5 minutes each

- Treat the tissue with Ir-Intercalator (Fluidigm) at a concentration of 1:400 in 1xPBS for 30 minutes

- Cleanse the slides in dH2O in dH2O for 5 minutes, leave to dry at least 20 minutes before running on the Hyperion

Ablation

- Ensure that the argon and helium gases are open and that the pressure is optimal (80kPa argon/50kPa helium)

- Start the ‘CyTOF’ software (Login: admin / no password), allow to initialise

- Click ‘Start’ in the top left corner, ensure the Plasma Start Up Sequence finishes successfully in the log manager (in the bottom left corner)

- Once the sequence is complete, leave for 30 minutes for the plasma to stabilise

- Load a manufacturer’s 3-Element Full Coverage Tuning Slide (Fluidigm) and run an ‘autotune’ 100µm wide and 500µm high, ensure all sub-calibrations are ticked before starting.

- Ensure that the ‘AutoTune’ has been successful by comparing the results to the word document ‘AutoTune for Calibration’ in the Desktop.

- Click on the ‘Data Acquisition’ tab and open a new .mcd file matching your slides specifications

- Take a picture of your slide, then right click on the template, and select import image. Ensure that the image is oriented with the slide label on the right-hand side.

  - We recommend taking a picture on your mobile device and emailing the image to your college email account, which can then be accessed via the IMC PC

- Use the panorama tool to draw an area you wish visualised, click on the panorama tab at the bottom of the screen and click ‘Start’ to acquire your image

- Once visualised draw an area you wish ablated using the ROI tool, click on the ROI tab at the bottom of the screen

- Ensure your ROI has a suitable template by right clicking the template box, either select a template or create a new one

  - A ‘template’ is a list of Ln-isotopes you wish to collate data for from the IMC detector

  - Select the isotopes required using the ‘Period Table’ tool

  - ROI will not start without a template, including ‘energy test’ ROI

- Ensure you have selected an energy level that cuts through the tissue but does not damage the glass below it

  - You may wish to run an ‘energy test’, draw several ROI (100µmx100µm)

  - Give them sequentially larger energy levels and click ‘Start’

  - The ROI that cuts through the tissue but doesn’t damage the glass is the best energy level for your tissue

    - Damage to the glass will appear as grid-like grooves in the slide

    - The laser should cut through as much tissue as possible, ideally leaving a clear ablation area with no obvious damage to the glass slide

- Once the ROI have been setup click ‘Start’ to begin ablation

- Once all tissue has been ablated shut-down the system by clicking ‘Stop’ in the top left corner of the screen

  - You may also tick the ‘Shutdown’ box in the ‘Data Acquisition’ tab to shut down the device when your final ROI completes

- Wait for the ‘Plasma Stop Sequence’ to finish successfully

- Turn off the argon and helium gas supply

Masking and Segmentation

- The IMC data will be saved as a .mcd file

- Open the data using the MCDViewer software available from Fluidigm

- Export your data by selecting File->Export, then select to export the data as either a single stacked ‘ome.tiff’ or several individual ‘ome.tiff’s for each channel

  - Either a single ‘ome.tiff’ stack or individual ‘ome.tiff’ files are viable options and can be opened in ImageJ

- Open these ‘ome.tiff’ using ImageJ (1.53c), then use the Despeckle function to remove single pixels using Process -> Noise -> Despeckle

- Select Image -> Adjust -> Brightness and Contrast and adjust to remove any background

- Save each individual channel as a ‘tiff’, keeping all .tiff files in sample specific folders

- Select 6 channels, plus 1 DNA channel, to create your mask. The channels are ideally cell markers denoting unique cell types and phenotypes.

- Change the colour of the channels, red for the cell markers and green for the nuclei, then save as a JPEG

- Open the Ilastik software (1.1.3post3) and start a ‘pixel classification’, then follow the steps of the Ilastik protocol, detailed below

  - Click ‘Add New. .’ and import your JPEG

  - Click ‘Feature Selection’ -> ‘Select Features…’ then tick every box in the ‘Features’ window

  - Click ‘Training’, then ‘add label’. Rename your labels nuclei, signal, and background, then start identifying each group. Only do a small section, then check your labelling by selecting ‘Live Update’ then ‘Segmentation’ and get an update on segmentation process

  - Click ‘Prediction Export’, then ‘Choose Export Image Settings…’ and ensure the format is set to ‘tiff’. Your JPEG can now be exported as a probability map for masking.

- Open the CellProfiler software (4.2.1) and open the ‘CellProfiler’ pipeline (details below), select ‘Analyse Images’. Masks will now be generated for your samples.

  - IdentifyPrimaryObjects will use the nuclei signal to create ‘objects’ each object representing a single nucleated cell

  - IdentifySecondaryObjects will expand around all primary objects (nuclei) with surrounding signal, associating all cell signals with a nucleus

  - IdentifyTeritaryObjects fills in the space between the primary and secondary objects, effectively representing the cytoplasm

  - MaskObjects will turn these objects into a ‘mask’ which can then be overlaid on the individual ‘.tiff’ channels to infer each cells characteristic

  - Each mask should be named after its corresponding sample and ROI

Data Analysis

- Ensure all individual channel ‘.tiff’ files and their associated masks in a sample are stored in distinct folders

- Open the HistoCAT software (1.76) then select Load -> Load Samples and import in the case folder containing all individual .tiff’s and the mask

- Accept the pixel expansion and do not transform data, then select a folder to save the ‘gates’. Please note this is where your data will be exported to.

- Once your files have loaded select ‘Save -> Export gates as CSV’, your data will now be accessible in a .csv format

- You may infer your ‘minimum positive value’, the lowest your cell signal can be in each channel that can be considered a positive signal, from your ImageJ edits previously the minimum set using the Brightness and Contrast tool.

  - Open the B&C tool then open the .tiff image, the minimum value will be represented on the B&C tool

  - You may also choose to record the minimum  values when you are first processing the images in ImageJ

- Colocalisation analysis can be done in one of three ways;

  - Pull the channel data from your .csv file, if both channels are positive there is colocalisation if one isn’t there is no colocalisation. This result can either be presented as a total number of colocalised cells or as a percentage of colocalised cells in each population.

  - Open the two channels on the ImageJ program, then use the ‘Analysis -> Colocalization -> Colocalization Threshold’, the level of colocalisation is generated as Rcoloc with a scale between 0 (no colocalisation) to 1 (complete colocalisation).

  - Merge the two channels together on the Image program, using ‘Image -> Color -> Merge Channels’, then manually count all the colocalised cells. . This result can either be presented as a total number of colocalised cells or as a percentage of colocalised cells in each population.

- Proximity analysis can be performed using the ‘Create ring regions new’ macro made by the PM Group.

  - Set the scale via ‘Analysis -> Set Scale’ and set ‘Distance in pixels’ to 1 and ‘Known distance’ to 1 and ‘Unit of Length’ to µm.

  - Then use the macro via ‘Plugin -> Macros -> Create ring  regions New’, open your channels and draw an ROI around the object being investigated. Add to the ROI manager (which should open when you start the macro).

  - Then open the channel of the signal you are investigating and run the macro and create 1 ring region 10µm thick. The average signal of the area around the object will now be recorded.

- Correlation analysis can be done by copying your channel data from the .csv file, then import the data into GraphPad prism and do a ‘Correlation Analysis’ to get an r value and a p value

- You can use the HistoCAT programme to perform neighbourhood clustering.

  - Select the samples to cluster and the channels of interest, then select the Analyse drop down menu in the bottom right of the screen and select ‘Run tSNE’

  - Two tSNE channels will be generated, highlight these channels, and select the Analyse drop down menu and select ‘Run PCA depending on your research question

  - Two new channels for the PCA will be generated, click the Analyse drop down menu and select ‘Scatter’ to display the clustered data

  - Select a channel from the ‘Heatmap Channels’ to display where each channel has been clustered.

  - All images in HistoCAT can be saved and exported via ‘Save -> Save Tiff Figure’

Accidental exposure/ first aid:

First Aid treatment for skin contact with xylene, ethanol, 10% neutral buffered formalin and deionized formamide:

Remove any contaminated clothing. Rinse skin well with water for a prolonged period. If necessary, seek medical attention.

First Aid treatment for inhalation:

Remove person to a well-ventilated area. If person seems to become dizzy or loses consciousness call security and wait for emergency services to arrive.

First Aid treatment for eye splash:

Rinse thoroughly for at least 15min using the eye wash station.  If necessary, seek medical attention.

First Aid treatment for ingestion:

Do NOT induce vomiting.  Rinse mouth with water and seek medical attention immediately.

First Aid treatment for scalds:

Flush the burn with lots of cool water for 20min.  Seek medical attention if the patient has anything more than a small minor burn.

Emergency support through security: 4444 (+442075891000)

In all instances of accident OR near miss, notify the safety department and complete a SALUS report.

https://www.imperial.ac.uk/safety/safety-by-topic/accidents--incidents/

Occupational health contacts:  https://www.imperial.ac.uk/occupational-health/

email: occhealth@imperial.ac.uk  phone: +44 20 7594 9401

| Date |  | Name |  | Signature |  | Trained by |  | Supervisor/Lab Manager |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Review history |  |  |  |  |  |  |  |  |  |
|  | Review 1 |  | Review 2 |  | Review 3 |  | Review 4 |  | Review 5 |
| Date conducted |  |  |  |  |  |  |  |  |  |
| Conducted by |  |  |  |  |  |  |  |  |  |